GPT-9900 Specifications

The specifications apply when the GPT-9900 is powered on for at least 30 minutes under +20°C~+30°C.

| Model \ Func. | AC WITHSTANDING | DC WITHSTANDING | INSULATION RESISITANCE | GROUND BOND |
| --- | --- | --- | --- | --- |
| GPT-9904 |  |  |  |  |
| GPT-9903 |  |  |  |  |
| GPT-9903A |  |  |  |  |
| GPT-9902A |  |  |  |  |
| GPT-9901A |  |  |  |  |

| # AC WITHSTANDING |  |  |  |
| --- | --- | --- | --- |
| Output-Voltage Range |  | 0.050kV~ 5.000kV ac |  |
| Output-Voltage Resolution |  | 2V/step |  |
| Output-Voltage Accuracy |  | ± (1% of setting +5V) [no load] |  |
| Maximum Rated Load |  | 500 VA (5kV/100mA) |  |
| Maximum Rated Current |  | 100mA (0.5kV < V ≦ 5kV) 10mA (0.05kV ≦ V≦ 0.5kV) |  |
| Output-Voltage Waveform |  | Sine wave |  |
| Output-Voltage Frequency |  | 50 Hz / 60 Hz selectable |  |
| Voltage regulation |  | ± (1% of reading +5V) [full load → no load] |  |
| Voltmeter Accuracy |  | ± (1% of reading+ 5V) |  |
| Current Measurement Range |  | 0.001mA~100.0mA |  |
| Current Best Resolution |  | 0.001mA / 0.01mA / 0.1mA |  |
| AC Current Measurement Accuracy |  | ±(1.5% of reading +30 counts) when HISET < 1.11mA ±(1.5% of reading +3 counts) when HISET ≧ 1.11mA |  |
| Window Comparator Method |  | Yes |  |
| ARC Detect |  | Yes |  |
| RAMP (Ramp-Up Time) |  | 0.1~999.9s |  |
| TIMER (Test Time)* |  | OFF, 0.5s~999.9s |  |
| Sweep Function* |  | Yes |  |
| GND |  | ON/OFF |  |
| # DC WITHSTANDING |  |  |  |
| Output-Voltage Range |  | 0.050kV~ 6.000kV dc |  |
| Output-Voltage Resolution |  | 2V/step |  |
| Output-Voltage Accuracy |  | ± (1% of setting +5V) [no load] |  |
| Maximum Rated Load |  | 100W (5kV/20mA) |  |
| Maximum Rated Current |  | 20mA (0.5kV < V ≦ 6kV) 2mA (0.05kV ≦ V≦ 0.5kV) |  |
| Voltage regulation |  | ± (1% of reading +5V) [full load → no load] |  |
| Voltmeter Accuracy |  | ± (1% of reading+ 5V) |  |
| Current Measurement Range |  | 0.001mA~020.0mA |  |
| Current Best Resolution |  | 0.001mA / 0.01mA / 0.1mA |  |
| DC Current Measurement Accuracy |  | ±(1.5% of reading +30 counts) when HISET < 1.11mA ±(1.5% of reading +3 counts) when HISET ≧ 1.11mA |  |
| Window Comparator Method |  | Yes |  |
| ARC Detect |  | Yes |  |
| RAMP (Ramp-Up Time) |  | 0.1~999.9s |  |
| TIMER (Test Time)* |  | OFF, 0.5s~999.9s |  |
| Sweep Function* |  | Yes |  |
| GND |  | ON/OFF |  |
| # INSULATION RESISITANCE |  |  |  |
| Output-Voltage |  | 50V~1000V dc |  |
| Output-Voltage Resolution |  | 50V/step |  |
| Output-Voltage Accuracy |  | ± (1% of setting +5V) [no load] |  |
| Resistance Measurement Range |  | 0.001 Gohm~ 50.00 Gohm |  |
|  | Test voltage | Measurable Range | Accuracy |
|  | 50V≦V≦450V | 0.001~ 0.050 Gohm 0.051~ 2.000 Gohm | ±(5% of reading +1 count) ±(10% of reading +1 counts) |
|  | 500V≦V≦1000V | 0.001~ 0.500 Gohm 0.501~ 9.999 Gohm 10.00~ 50.00 Gohm | ±(5% of reading +1 count) ±(10% of reading +1 count) ±(20% of reading +1 count) |
| Window Comparator Method |  | Yes |  |
| Output Impedance |  | 600 kohm |  |
| RAMP (Ramp-Up Time) |  | 0.1~999.9s |  |
| TIMER (Test Time) |  | 1s~999.9s |  |
| Sweep Function* |  | Yes |  |
| GND |  | OFF (fixed) |  |
| ### GROUND BOND |  |  |  |
| Output-Current Range |  | 03.00A~32.00A ac |  |
| Output-Current Resolution |  | 0.01A |  |
| Output-Current Accuracy |  | 3A≦I≦8A : ± (1% of setting +0.2A) 8A<I≦32A : ± (1% of setting +0.05A) |  |
| Test-Voltage |  | 6Vac max. (open circuit) |  |
| Test-Voltage Frequency |  | 50 Hz / 60 Hz selectable |  |
| Resistance Measurement Range |  | 10mW~650.0m ohm |  |
| Resistance Measurement Resolution |  | 0.1mW |  |
| Resistance Measurement Accuracy |  | ± (1% of reading +2m ohm) |  |
| Window Comparator Method |  | Yes |  |
| TIMER (Test Time) |  | 0.5s~999.9s |  |
| Sweep Function* |  | Yes |  |
| GND |  | OFF (fixed) |  |
| Test Method |  | Four Terminal |  |
| ## MEMORY |  |  |  |
| Single Step Memory |  | MANU : 100 blocks |  |
| Automatic Testing Memory |  | AUTO : 100 blocks, menu per auto : 16 |  |
| ## INTERFACE |  |  |  |
| Rear Output |  | Standard |  |
| RS232 |  | Standard |  |
| USB (Device) |  | Standard |  |
| GPIB |  | Option |  |
| Remote Terminal (Front) |  | Standard |  |
| SIGNAL IO |  | Standard |  |
| ## General |  |  |  |
| DISPLY |  | 240 x 64 dot matrix LED back light LCD |  |
| POWER SOURCE |  | AC100V/120V/220V/230V ±10%, 50/60Hz |  |
| POWER CONSUMPTION |  | 1000VA max. |  |
| DIMENSIONS & WEIGHT |  | 330(W) x 148(H) x 587(D) mm (GPT-9904/9903), 330(W) x 148(H) x 482(D) mm (GPT-9903A/9902A/9901A) approx.27kg(Max) |  |

* The “Sweep” and “TIMER OFF” can only be performed when the tester is in the special MANU mode